| Version History |  |  |
| --- | --- | --- |
| Version | Description of Change | Change Author |
| 2.0 | Initial release. | Venkata siva prasad |

Table of Contents

1.	Executive Summary	3

1.2 Accessory Products	4

1.3 Summary of Results	4

2.	Applicable Documents	17

2.1	References	17

2.2	Acronyms & Abbreviations	18

3.	System Overview	18

4 	Test Methods	19

5.	Sample Size	19

6.	Sample Preconditioning	20

7.	Sample Maturity	20

8.	Deviation	20

9.	 Equipment And Non-Product Software	20

10.	Test Sample Retention	20

11.	Raw Data Retention	20

12. 	Conclusions	20

# Executive Summary

The scope of this system regression was to assess the overall system functional performance of the updated release candidate of clinician communicator (8880T2) and A901 Communication Manager. Backward compatibility is tested with the Restore therapy fielded application of Spinal Cord Stimulation SCS devices (PrimeAdvancedMRI™, RestoreUltraMRI™, TestStim, RestoreSensorMRI™) with the tablet Samsung S2 running OS 7.0, Samsung S5e running OS 9.0 and with MyStim Patient Programmer (97740).

The overall system functional performance of the updated releases of A4P tablet with OS14, clinician communicator (8880T2), Restore therapy application, PrimeAdvancedMRI™ INS (SCS), communication manager and Patient Data Service. The test and handset (A10e) (Refer versions in Section 1.1) of Spinal Cord Stimulation (SCS). The updated versions are tested with Patient Programmer A10e (Refer versions in Section 1.1).

This report demonstrates that all test results observed during regression testing matched the expected test results specified in the SCS System Regression Test Protocol – NDHF1579-214568. The objective evidence gathered when executing said protocol demonstrates that the products specified in the “Products Under Test Section 1.1” table below continue to work as intended.

1.1 Products Under Test

The following products were subjected to regression testing as outlined in SCS System Regression Test Protocol – NDHF1579-214568 (Section 13 Test Case 6):

| Model # | Description | Version Tested |
| --- | --- | --- |
| 97740 | MyStim Patient Programmer | 9.3 |
| SM-T813 | Samsung Galaxy S2 | OS 7.0 |
| SM-T720 | Samsung Galaxy S5E | OS 9.0 |
| SM-T630 | Samsung Galaxy A4P | OS 14 |
| A71100 | “Restore” Clinician Programmer S2, S5e | 1.0.4318 |
| A71100 | “Restore” Clinician Programmer A4P | 1.0.4478 |
| A901 | Communications Manager Application S2, S5e | 1.0.1325 |
| A901 | Communications Manager Application A4P | 2.0.1367 |
| A902 | Patient Data Services Application S2 | 1.0.828 |
| A902 | Patient Data Services Application S5e | 1.0.929 |
| A902 | Patient Data Services Application A4P | 1.0.947 |
| 8880T2 | Clinician Communicator (CTM) | 1.123 |

## 1.2 Accessory Products

The following products were utilized to conduct system regression testing, but were not evaluated for impact to previously completed system design verification testing:

| Model # | Description |
| --- | --- |
| 97702 | Restore Prime Advanced SureScan MRI Neurostimulator |
| 97712 | Restore Ultra Sure Scan MRI Neurostimulator |
| 37022 | Test Stim Neurostimulator |
| 355531 | Multi Lead Trialing Cable |
| 97714 | Restore Sensor SureScan MRI Neurostimulator |
| 8880T2 | Communicator |
| 8840T | Telemetry head |

## 1.3 Summary of Results

All regression test cases met expected criteria and thus the products under test continue to work as intended.

RestoreUltraMRITM

| Test Case Executed | Test Case Description | Test Operator/Test Date | Result |
| --- | --- | --- | --- |
| - Spinal Cord Stimulation - Test Case 1.12: RestoreUltraMRI Test configuration with LED & measured Amplitude. | Amplitudes measured on 16 electrodes for legacy therapy groups using LED Box. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Amplitude LEDs illuminated for each configured electrode as expected. Pass |
| Spinal Cord Stimulation - Test Case 1.14: Change Stimulation state from ON to OFF | Stimulation from ON to OFF is checked using LED box. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Stimulation state changed from ON to OFF. Pass |
| Spinal Cord Stimulation - Test Case 1.15: Impedance Classification Clinician Programmer App UI. | Checked for Electrode impedance using clinician application. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Electrode impedances that were set using the impedance box are classified as “Good” (>50Ω and <10,000Ω) in the Clinician programmer app. Pass |
| Spinal Cord Stimulation - Test Case 1.21: Session Report observed parameters. | Record the following parameters from session report: - Patient First Name - Patient Last Name - Device Model # - Lead Model Numbers - Program(s) at End Of Session - Electrode Impedance | S2: Venkata Hiranmai, 614848/1 August 2024. (Electrode Impedance Only). S5e: Venkata Hiranmai, 614848/2 August 2024. | The following parameters are observed to be present in session report: - Patient First Name - Patient Last Name - Device Model # - Lead Model Numbers - Program(s) at End Of Session - Electrode Impedance Pass |
| Spinal Cord Stimulation - Test Case 1.23: MRI Eligibility. | Record the observed MRI Eligibility on Clinician Programmer App UI and the MRI report | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Observed that neurostimulation system MRI scan-type eligibility cannot be determined. Pass |
| Spinal Cord Stimulation - Test Case 1.25: Stimulation state OFF to ON using MyStim Patient Programmer. | Testing stimulation from OFF to ON is checked using LED box. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Stimulation turned OFF to ON Pass |
| Spinal Cord Stimulation - Test Case 1.26: Adjust Amplitude | Check that the Patient Programmer allows amplitude to be incremented and decremented | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Patient programmer app allows amplitude to be incremented and decremented as allowed by clinician app. Pass |
| Spinal Cord Stimulation - Test Case 1.27: Adjust Pulse Width | Check that the Patient Programmer allows pulse width to be incremented and decremented | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Patient programmer app allows pulse width to be incremented and decremented as allowed by clinician app. Pass |
| Spinal Cord Stimulation - Test Case 1.28: Adjust Rate. | Check that the Patient Programmer allows Rate to be incremented and decremented | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Patient programmer app allows Rate to be incremented and decremented as allowed by clinician app. Pass |
| Spinal Cord Stimulation - Test Case 1.29: Stimulation state ON to OFF using MyStim Patient Programmer. | Check that the therapy is turned OFF using patient programmer and verify on LED box. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Stimulation turned ON to OFF. Pass |

TestStim

| Test Case Executed | Test Case Description | Test Operator/Test Date | Result |
| --- | --- | --- | --- |
| - Spinal Cord Stimulation - Test Case 1.12: TestStim Test configuration with LED & measured Amplitude. | Amplitudes measured on 16 electrodes for legacy therapy groups using LED Box. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Amplitude LEDs illuminated for each configured electrode as expected. Pass |
| Spinal Cord Stimulation - Test Case 1.14: Change Stimulation state from ON to OFF | Stimulation from ON to OFF is checked using LED box. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Stimulation state changed from ON to OFF. Pass |
| Spinal Cord Stimulation - Test Case 1.15: Impedance Classification Clinician Programmer App UI. | Checked for Electrode impedance using clinician application. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Electrode impedances that were set using the impedance box are classified as “Good” (>50Ω and <10,000Ω) in the Clinician programmer app. Pass |
| Spinal Cord Stimulation - Test Case 1.21: Session Report observed parameters. | Record the following parameters from session report: - Patient First Name - Patient Last Name - Device Model # - Lead Model Numbers - Program(s) at End Of Session - Electrode Impedance | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | The following parameters are observed to be present in session report: - Patient First Name - Patient Last Name - Device Model # - Lead Model Numbers - Program(s) at End Of Session - Electrode Impedance Pass |
| Spinal Cord Stimulation - Test Case 1.25: Stimulation state OFF to ON using MyStim patient programmer. | Testing stimulation from OFF to ON is checked using LED box. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Stimulation turned OFF to ON Pass |
| Spinal Cord Stimulation - Test Case 1.26: Adjust Amplitude | Check that the Patient Programmer allows amplitude to be incremented and decremented | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Patient programmer app allows amplitude to be incremented and decremented as allowed by clinician app. Pass |
| Spinal Cord Stimulation - Test Case 1.27: Adjust Pulse Width | Check that the Patient Programmer allows pulse width to be incremented and decremented | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Patient programmer app allows pulse width to be incremented and decremented as allowed by clinician app. Pass |
| Spinal Cord Stimulation - Test Case 1.28: Adjust Rate. | Check that the Patient Programmer allows Rate to be incremented and decremented | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Patient programmer app allows Rate to be incremented and decremented as allowed by clinician app. Pass |
| Spinal Cord Stimulation - Test Case 1.29: Stimulation state ON to OFF using MyStim Patient Programmer. | Check that the therapy is turned OFF using patient programmer and verify on LED box. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Stimulation turned ON to OFF Pass |

RestoreSensorMRITM

| Test Case Executed | Test Case Description | Test Operator/Test Date | Result |
| --- | --- | --- | --- |
| - Spinal Cord Stimulation - Test Case 1.12: RestoreSensorMRI Test configuration with LED & measured Amplitude. | Amplitudes measured on 16 electrodes for legacy therapy groups using LED Box. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Amplitude LEDs illuminated for each configured electrode as expected. Pass |
| Spinal Cord Stimulation - Test Case 1.14: Change Stimulation state from ON to OFF | Stimulation from ON to OFF is checked using LED box. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Stimulation state changed from ON to OFF. Pass |
| Spinal Cord Stimulation - Test Case 1.15: Impedance Classification Clinician Programmer App UI. | Checked for Electrode impedance using clinician application. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Electrode impedances that were set using the impedance box are classified as “Good” (>50Ω and <10,000Ω) in the Clinician programmer app. Pass |
| Spinal Cord Stimulation - Test Case 1.21: Session Report observed parameters. | Record the following parameters from session report: - Patient First Name - Patient Last Name - Device Model # - Lead Model Numbers - Program(s) at End Of Session - Electrode Impedance | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | The following parameters are observed to be present in session report: - Patient First Name - Patient Last Name - Device Model # - Lead Model Numbers - Program(s) at End Of Session - Electrode Impedance Pass |
| Spinal Cord Stimulation - Test Case 1.23: MRI Eligibility. | Record the observed MRI Eligibility on Clinician Programmer App UI and the MRI report | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Observed that neurostimulation system MRI scan-type eligibility cannot be determined. Pass |
| Spinal Cord Stimulation - Test Case 1.25: Stimulation state OFF to ON using myStim Patient Programmer. | Testing stimulation from OFF to ON is checked using LED box. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Stimulation turned OFF to ON Pass |
| Spinal Cord Stimulation - Test Case 1.26: Adjust Amplitude | Check that the Patient Programmer allows amplitude to be incremented and decremented | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Patient programmer app allows amplitude to be incremented and decremented as allowed by clinician app. Pass |
| Spinal Cord Stimulation - Test Case 1.27: Adjust Pulse Width | Check that the Patient Programmer allows pulse width to be incremented and decremented | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Patient programmer app allows pulse width to be incremented and decremented as allowed by clinician app. Pass |
| Spinal Cord Stimulation - Test Case 1.28: Adjust Rate. | Check that the Patient Programmer allows Rate to be incremented and decremented | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Patient programmer app allows Rate to be incremented and decremented as allowed by clinician app. Pass |
| Spinal Cord Stimulation - Test Case 1.29: Stimulation state ON to OFF using MyStim Patient Programmer. | Check that the therapy is turned OFF using patient programmer and verify on LED box. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Stimulation turned ON to OFF Pass |

Restore PrimeAdvancedMRI™

| Test Case Executed | Test Case Description | Test Operator/Test Date | Result |
| --- | --- | --- | --- |
| - Spinal Cord Stimulation - Test Case 1.12: PrimeAdvancedMRI Test configuration with LED & measured Amplitude. | Amplitudes measured on 16 electrodes for legacy therapy groups using LED Box. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Amplitude LEDs illuminated for each configured electrode as expected. Pass |
| Spinal Cord Stimulation - Test Case 1.14: Change Stimulation state from ON to OFF | Stimulation from ON to OFF is checked using LED box. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Stimulation state changed from ON to OFF. Pass |
| Spinal Cord Stimulation - Test Case 1.15: Impedance Classification Clinician Programmer App UI. | Checked for Electrode impedance using clinician application. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Electrode impedances that were set using the impedance box are classified as “Good” (>50Ω and <10,000Ω) in the Clinician programmer app. Pass |
| Spinal Cord Stimulation - Test Case 1.21: Session Report observed parameters. | Record the following parameters from session report: - Patient First Name - Patient Last Name - Device Model # - Lead Model Numbers - Program(s) at End Of Session - Electrode Impedance | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | The following parameters are observed to be present in session report: - Patient First Name - Patient Last Name - Device Model # - Lead Model Numbers - Program(s) at End Of Session - Electrode Impedance Pass |
| Spinal Cord Stimulation - Test Case 1.23: MRI Eligibility. | Record the observed MRI Eligibility on Clinician Programmer App UI and the MRI report | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Observed that neurostimulation system MRI scan-type eligibility cannot be determined. Pass |
| Spinal Cord Stimulation - Test Case 1.25: Stimulation state OFF to ON using MyStim Patient Programmer. | Testing stimulation from OFF to ON is checked using LED box. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Stimulation turned OFF to ON Pass |
| Spinal Cord Stimulation - Test Case 1.26: Adjust Amplitude | Check that the Patient Programmer allows amplitude to be incremented and decremented | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Patient programmer 97740 allows amplitude to be incremented and decremented as allowed by clinician app. Pass |
| Spinal Cord Stimulation - Test Case 1.27: Adjust Pulse Width | Check that the Patient Programmer allows pulse width to be incremented and decremented | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Patient programmer 97740 allows pulse width to be incremented and decremented as allowed by clinician app. Pass |
| Spinal Cord Stimulation - Test Case 1.28: Adjust Rate. | Check that the Patient Programmer allows Rate to be incremented and decremented | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Patient programmer 97740 allows Rate to be incremented and decremented as allowed by clinician app. Pass |
| Spinal Cord Stimulation - Test Case 1.29: Stimulation state ON to OFF using MyStim Patient Programmer. | Check that the therapy is turned OFF using patient programmer and verify on LED box. | S2: Venkata Hiranmai, 614848/1 August 2024. S5e: Venkata Hiranmai, 614848/2 August 2024. | Stimulation turned ON to OFF. Pass |

# Applicable Documents

## References

| Document Number | Version | Document Title |
| --- | --- | --- |
| QMS16001 | G | Test Method Validation |
| QMSWI16301 | A | Attribute Test Method Validation |
| PCPWI33003 | E | Design Capability Assessment |
| NSP1009-22797 | 3.0 | Validation Report for Rx1 Lab Programmer |
| NDHF1245-127213 | 2.0 | Validation Plan for INS Reset Script |
| NSP1009-30347 | 2.0 | Source Archive, Rx1 Lab Programmer |
| NDHF1245-141451 | 2.0 | RS2 Lab Programmer Release 3.3.0 Code |
| NDHF1579-175767 | 15.0 | System Regression Testing Protocol - Cross Therapies |
| NDHF1245-108853 | 36.0 | RS2 System Interface and Architecture |
| NDHF1579-228311 | 18.0 | SCS Restore Regression Raw Data for S2, S5e, A4P |
| NDHF1579-228423 | 2.0 | Signature Page for NDHF1579-228311 |
| NDHF1579-214568 | 4.0 | SCS System Regression Test Protocol |
| QMS23875 | D | Neuromodulation Calibration Procedure |
| QMS23873 | D | Neuromodulation Equipment Qualification Procedure |

## Acronyms & Abbreviations

| Term | Definition |
| --- | --- |
| BT | Bluetooth |
| COTS | Commercial off-the-shelf |
| CP App | Clinician Programmer therapy software application |
| CTM | Clinician Telemetry Module |
| ELN | Electronic Lab Notebook |
| FRB | Functional Review Board |
| INS | Implantable Neuro Stimulator |
| PDS | Patient Data Services |
| PP App | Patient Programmer Application |
| PTM | Patient Therapy Manager |
| SCS | Spinal Cord Stimulation |
| TM | Telemetry Module (a.k.a. Fob) |

# System Overview

The Restore is designed to meet the SCS therapy system need to deliver clinically adjustable electrical stimulation pulses through electrodes to targeted regions of the spinal cord. Peripheral Nerve Stimulation (PNS) is a palliative medical therapy system designed to deliver clinically adjustable electrical stimulation pulses through electrodes to targeted peripheral nerves outside of the central nervous system. A system that supports both SCS and PNS therapies is commonly (and for the purposes of this document) referred to as an “SCS system.” The RS2 system will use the new Clinician Programmer (CP), Clinician Telemetry Module (CTM) and Patient Therapy Manager (PTM).

This electrical stimulation consists of a continuous series of electrical pulses that are generated by a battery powered Implantable Neurostimulator (INS). These electrical pulses are transmitted to the spinal cord target location using electromechanically connected therapy delivery elements.

At the distal end of the system are a set of tissue-exposed metal electrodes. Electrical stimulation from the INS causes ionic current to flow through the tissue to the targeted spinal cord nerve bundles by inducing a redox (reduction-oxidation) reaction at the electrode-tissue interface. This ionic current modulates the physiological transmission of neural axon action potentials within the spinal cord, thus evoking a therapeutic response.

The primary purposes of SCS and PNS are to improve patient quality-of-life and physical function by reducing the severity of pain and its concomitant effects. Afferent pain signals in dorsal tracts of the spinal cord are inhibited by SCS therapy before they reach the brain, providing a degree of relief from chronic pain.

Figure 1: System Architecture Diagram

Restore INS

The INS developed for the Restore System uses several building blocks and mechanical design features (mechanical form factor).

Common Programmer Platform SW

- Communication Manager Application (Model A901) – Contains the communications functionality to enable data transmission between the therapy application and the neurostimulator device.

- Patient Data Services Application (Model A902) – Houses a database

Communicator – Next-generation telemetry module.

MyStim Patient Programmer

Restore devices uses MyStim Patient Programmer. Mystim system element wirelessly communicates with INS via proprietary telemetry protocols.

# 4 	Test Methods

A portion of the testing detailed in this regression testing report was conducted utilizing LED Box.

# 5.	Sample Size

All testing completed in this regression protocol are invariant through multiple runs, due to its purely software read/write nature. In these scenarios, variability does not exist; therefore, per PCPWI33003, a sample size of 1 is sufficient for demonstrating capability for these regression testing.

# 6.	Sample Preconditioning

Given that the nature of this regression is software centric, no sample preconditioning is needed. The handset was preconfigured by AirWatch under Group ID RPECPS2ID / RPECPSEID and relevant software application information is documented in NDHF1579-228311.

# 7.	Sample Maturity

Hardware will be production equivalent and software will be at least at Release Candidate. Traceability of model tested, and operating system will be recorded to show which configuration was tested.

# 8.	Deviation

No Deviations.

# 9.	 Equipment And Non-Product Software

Equipment utilized was calibrated per Neuromodulation Calibration Procedure (QMS23875) and qualified per Neuromodulation Equipment Qualification (QMS23873) prior to each test configuration execution

# 10.	Test Sample Retention

No sample retention is required given that the nature of this regression testing is software centric. The hardware is production equivalent and software are version controlled which can be used to re-configure the hardware, if needed.

# 11.	Raw Data Retention

Raw data collected in NDHF1579-228311, SCS Restore Regression Raw Data for S2, S5e, A4P.

# 12. 	Conclusions

All observed test results matched expected test results, thus demonstrating the products specified in the “Products Under Test” table in the Executive Summary above, continue to work as intended.

END OF DOCUMENT