## Test Case 1: Restore Clinician and MyStim Patient Programmers

Test Case 1: Medtronic Products

- Restore Clinician Programmer Application, Model # A71100

- Communication Manager Application, Model # A901

- Patient Data Services Application, Model # A902

- Mobile Platform – Tablet, Model # CT900

- Communicator, Model # 8880 T2

- Itrel 4B Neurostimulator, Model # 37704

- Restore Ultra Neurostimulator, Model # 37712 OR Restore Ultra Sure Scan MRI Neurostimulator, Model #97712

- Test Stim Neurostimulator, Model # 37022

- Multi Lead Trialing Cable, Model 355531

- Restore Sensor SureScan MRI Neurostimulator, Model # 97714

- Restore Prime Advanced Neurostimulator, Model # 37702 OR Restore Prime Advanced SureScan MRI Neurostimulator, Model #97702

- Restore Neurostimulator, Model # 37711 OR Restore Ultra SureScan MRI Neurostimulator, Model #97712

- Restore Advanced SureScan MRI Neurostimulator, Model #97713

- MyStim Patient Programmer, Model# 97740

Test Case 1: Equipment and Non-Product Software

- LED Stimulation Output Box with DIP switch for Can connection.

- Oscilloscope and Differential Probe

- Multimeter

- 8840T Telemetry Head (for use with Lab Programmer)

- TMTI Telemetry-M Test Interface (for use with Lab Programmer)Medtronic Credentialed Computer with Rx1 Lab Programmer installed (See NSP1009-22797 for Validation Report)

Test Case 1: Test Results, Operator(s), Date(s), and Data Traceability

Test results, operator(s), date(s), and data traceability will be published in the Report.

Test Case 1: Initial Test Set-Up

- Using NDHF1465-180952 Pulse Measurement Tool Test Method, sections 5.1 through 5.3, set up the Variable LED Display Box, differential amplifier and oscilloscope.

  - Use Option 1 or Option 2 of section 5.3.1 for interfacing with the lead.

  - Apply load to the electrodes per Option 1 or Option 2 per section 5.3.1 using incorporated DIP switches. If the LED Stimulation Display Box has low-z leads (< 20 Ω), set the DIP switches to “010” for 500 Ω.  Otherwise, for high-z leads (> 100 Ω), set the DIP switches to “100” for 300 Ω. Set the Can DIP switches to “111” for 150 Ω.

- Record the LED Stimulation Output Box, Oscilloscope, Differential Probe and Multimeter information below. Set each Neurostimulator to be used for test, as identified in the Restore Test Configuration Table, to shelf state following Restore Testing – Set Neurostimulators to Shelf State.  Record the 8840T Telemetry Head and Telemetry-M Test information below prior to launching the lab programmer software.

Restore Equipment and Non-Product Software List

| Equipment Name | Model or PN | Serial # or ID # or Version | Equipment Qual or Validation Report # | Calibration Due Date |
| --- | --- | --- | --- | --- |
| LED Stimulation Display Box | N/A | Box M1 | NRP1025-55414 | No Calibration Required |
| Oscilloscope | N/A | N/A | N/A | N/A |
| Differential Probe | N/A | N/A | N/A | N/A |
| Multimeter | N/A | N/A | N/A | N/A |
| Telemetry Head | 8840T | NHG413210N | No Calibration Required | No Calibration Required |
| TMTI Telemetry-M Test Interface | TMTI-11-RC-F-D-P | 11C0592 | No Calibration Required | No Calibration Required |
| Rx1 Lab Programmer | NSP1009-30347 | 5.4.1 | NSP1009-22797 | No Calibration Required |

- Record the Tablet Serial # Android version (from device setting) in the Restore Medtronic Product List.

- Record the MyStim Patient Programmer and Version Number in the Restore Medtronic Product List.

  - The MyStim serial number may be found inside of the battery compartment by removing batteries.

  - The version of the MyStim Patient programmer may be determined by removing and then re-inserting batteries.  After re-inserting the batteries, the version will appear in the lower left hand corner of the MyStim screen.

Restore Medtronic Product List

| Product Name | Model # | Serial # / Lot # | Version # |
| --- | --- | --- | --- |
| Restore Clinician Programmer Application | A71100 | N/A | Version: 1.0.4488 |
| Communication Manager Application | A901 | N/A | Version: 1.0.1325 |
| Patient Data Service Application | A902 | N/A | Version: 1.0.947 |
| Mobile Platform – Tablet | CT900E, SM-T720 | Serial Number: R52NB09DX5M | Navigate to Tablet Settings -> About tablet -> Software information to record the following information: Android Version: 9 Android Security Patch level: April 1, 2020 Build Number: PPR1.180610.011.T720XXU1ATE3 Navigate to Hub application in the tablet to record the following information: AirWatch Group ID: RPECPS5eid |
| Communicator | 8880T2 | Serial Number: NKW026179N | FW App Version: 1.123 |
| MyStim Patient Programmer | 97740 | Serial Number: NMH503177N | Version: 9.3 |
| Restore Ultra Neurostimulator | 37712 / 97712 | Model Number: 97712 Serial Number: NMB708792 | FW App ID: 01.03 FW App Version: 02.32 |
| Test Stim Neurostimulator | 37022 | Serial Number: NJS018943 | FW App ID: 01.03 FW App Version: 02.16 |
| Multi Lead Trialing Cable | 355531 | Lot Number: N756361 | N/A |
| Restore Sensor SureScan MRI Neurostimulator | 97714 | Serial Number: NMD762983 | FW App ID: 01.05 FW App Version: 02.02 |
| Restore Prime Advanced Neurostimulator | 37702 / 97702 | Model Number: 97702 Serial Number: NMA721252 | FW App ID: 01.02 FW App Version: 02.03 |
| Restore Neurostimulator | 37711 / 97712 | Model Number: N/A Serial Number: N/A | FW App ID: N/A FW App Version: N/A |
| Itrel 4B Neurostimulator | 37704 | Serial Number: NMY701644 | FW App ID: 01.04 FW App Version: 02.20 |
| Restore Advanced SureScan MRI Neurostimulator | 97713 | Serial Number: NMC702965 | FW App ID: 01.01 FW App Version: 04.07 |

Restore Test Configuration Table

| Test Config | Neurostimulator | Star Load Connection Method | Connected Electrodes | Torque Wrench |
| --- | --- | --- | --- | --- |
| 1 | Restore Ultra Neurostimulator (37712/97712) | LED Box internally connected 1x8 SCS Leads (2) | E0-E15 | 5 in-oz (White) |
| 2 | Restore Test Stim Neurostimulator (37022) | LED Box internally connected 1x8 SCS Leads (2) and Multi Lead Trialing Cable | E0-E15 | N/A |
| 3 | Restore Sensor SureScan MRI Neurostimulator (97714) | LED Box internally connected 1x8 SCS Leads (2) | E0-E15 | 4 in-oz (White or Blue) |
| 4 | Restore Prime Advanced Neurostimulator (37702 / 97702) | LED Box internally connected 1x8 SCS Leads (2) | E0-E15 | 5 in-oz (White) |
| 5 | Restore Neurostimulator (37711 / 97712) | LED Box internally connected 1x8 SCS Leads (1) | E0-E7 | 5 in-oz (White) |
| 6 | Itrel 4B Neurostimulator (37704) | LED Box internally connected 1x8 SCS Leads (1) | E0-E4, Case | 4 in-oz (White or Blue) |
| 7 | RestoreAdvancedMRI Neurostimulator (97713) | LED Box internally connected 1x8 SCS Leads (2) | E0-E15 | 5 in-oz (White) |

Restore Therapy Configuration Table

| Test Config | Lead Model (Bore) | Implant Location | Tip Location | Abandon Leads, Extensions | Group | Program | Cathode (-) | Anode (+) | Program Amplitude (V) | Program Pulse Width (μs) | Group Rate (Hz) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 977A1 (1) 977A2 (2) | Buttocks (Right) | Both Cervical | No Extensions | A | 1 | E0, E6, E10, E12 | E1, E7, E11, E13 | 2.50 | 200 | 160 |
| 2 | 3777 (1) 3877 (2) | N/A | N/A | No Extensions | B | 2 | E1, E7, E11, E13 | E0, E6, E10, E12 | 5.00 | 80 | 400 |
| 3 | 3873 (1) 3778 (2) | Buttocks (Right) | Both Cervical | Both No | C | 3 | E2, E4, E8, E14 | E3, E5, E9, E15 | 7.40 | 90 | 800 |
| 4 | 39286 (Drop 1) * | Buttocks (Right) | N/A | No Extensions | C | 4 | E3, E5, E9, E15 | E2, E4, E8, E14 | 9.55 | 120 | 80 |
| 5 | 3888/3892 (1) | Buttocks (Right) | Cervical | No Extensions | B | 2 | E0 | E3 | 10.00 | 70 | 65 |
| 6 | 3487A/3890 (1) | Buttocks (Right) | N/A | N/A | A | 1 | E0, E1, E2, E3 | C | 6.10 | 60 | 60 |
| 7 | 977A1 (1) 977A2 (2) | Buttocks (Right) | Both Cervical | No Extensions | A | 1 | E0, E6, E10, E12 | E1, E7, E11, E13 | 2.50 | 200 | 120 |

Note(s):

- Bore 1 corresponds to electrodes E0-E7 (For non-Itrel 4 Neurostimulators, E0-E3 for Itrel 4 Neurostimulator).

- When the lead compatibility is not under test, use any lead model to make electrical connection with Neurostimulator.

- Bore 2 corresponds to E8-E15.

- *Lead Model 39826 is a single, 16 contact, surgical lead.  The application allows this lead to be dropped on Bore 1 or Bore 2. For purposes of this testing, it is required to be dropped on Bore 1 to keep a 1:1 header contact to LED Stimulation Output Box Electrode labeling relationship.

Test Case 1: Test Procedure for each Test Configuration

- Connect the Neurostimulator to the LED Box. Except for the Itrel 4, attach the LED Box internally connected 1x8 Leads corresponding to E0-E7 to the Neurostimulator bore associated to E0-E7 and the LED Box internally connected 1x8 leads corresponding to E8-E15 to the Neurostimulator bore associated to E8-E15. For Itrel 4B Neurostimulator (37704):

  - Attach the LED Box internally connected 1x8 Leads corresponding to E0-E7 to the Neurostimulator bore associated to E0-E3.

  - Clasp a case clip around the Itrel 4B case. Attach this clip to the ‘’Can/Gnd” Clip/DIP switch. Set the DIP Switch associated to “Can” to “111” (corresponding to 150 Ω).

  - Set all other DIP switches, E4-E15, to “000” so they are isolated from the circuit.

- Connect to the Neurostimulator with the Restore Clinician Programmer Application. Select START WARRANTY or START USAGE and then OK to start device usage. Select START under IMPLANT DEVICE or START EVALUATION and then OK.

- Select Implant Location as defined in the Restore Therapy Configuration Table and then hit UPDATE.

- Use the left navigation drawer to go to About System and record the Neurostimulator’s Serial Number, FW App ID, FW App Version in the Restore Medtronic Product List above.

- For the first test configuration only, record the Model A71100 Restore Clinician App Version, Communication Manager Version, and Patient Data Service Version within the Tablet tab in the Restore Medtronic Product List above.

- For the first test configuration only, record the 8880T2 Communicator Serial Number and FW App Version within the Communicator tab in the Restore Medtronic Product List above.

- Close “About System” and navigate to Patient Info screen and select any Diagnosis.

- Enter current Test Config #, as defined in the Restore Test Configuration Table, into FIRST NAME field.

- Enter “Regression” into LAST NAME field and then hit UPDATE.

- Go to the Lead Select screen and configure the leads as defined in the Restore Therapy Configuration Table and then hit UPDATE. If an MRI SCAN ELIGIBILITY dialog is displayed, hit UPDATE to proceed.

- Set the stimulation configuration parameters (electrode configuration, Amplitude, Pulse Width, Group Rate) as defined in the Restore Therapy Configuration Table.  When prompted, hit STIM ON.

- Record observed Polarity LED State / Color as well as number of Amplitude LEDs illuminated for each configured electrode below in the correct test config #.  Note that the number of Amplitude LEDs illuminated has no expected results in this step, but it is used as expected results in step 26.  Verify the observed output matches the expected output. Indicate Pass or Fail for each.  Referring to NDHF1465-180952, Section 5.9, the Polarity LED State and Color may be classified as follows:

| Electrode State | Polarity LED State / Color |
| --- | --- |
| Cathode (-) | Illuminated, Red |
| Anode (+) | Illuminated, Green |
| Unassigned | Not Illuminated, None |

Test Config #1

| Electrode | Polarity LED State / Color (Expected) | Polarity LED State / Color (Actual) | # of Amplitude LEDs Illuminated (Actual) | Polarity LED State / Color (Pass /Fail) |
| --- | --- | --- | --- | --- |
| E0 | Illuminated, Red | Illuminated, Red | 1 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E1 | Illuminated, Green | Illuminated, Green | 1 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E2 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E3 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E4 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E5 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E6 | Illuminated, Red | Illuminated, Red | 1 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E7 | Illuminated, Green | Illuminated, Green | 1 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E8 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E9 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E10 | Illuminated, Red | Illuminated, Red | 1 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E11 | Illuminated, Green | Illuminated, Green | 1 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E12 | Illuminated, Red | Illuminated, Red | 1 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E13 | Illuminated, Green | Illuminated, Green | 1 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E14 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E15 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| Refer RestoreUltraMRI_CP_Stim_ON_LED_Test.jpg |  |  |  |  |

Test Config #2

| Electrode | Polarity LED State / Color (Expected) | Polarity LED State / Color (Actual) | # of Amplitude LEDs Illuminated (Actual) | Polarity LED State / Color (Pass /Fail) |
| --- | --- | --- | --- | --- |
| E0 | Illuminated, Green | Illuminated, Green | 3 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E1 | Illuminated, Red | Illuminated, Red | 3 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E2 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E3 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E4 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E5 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E6 | Illuminated, Green | Illuminated, Green | 3 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E7 | Illuminated, Red | Illuminated, Red | 3 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E8 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E9 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E10 | Illuminated, Green | Illuminated, Green | 3 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E11 | Illuminated, Red | Illuminated, Red | 3 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E12 | Illuminated, Green | Illuminated, Green | 3 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E13 | Illuminated, Red | Illuminated, Red | 3 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E14 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E15 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| Refer TestStim_CP_Stim_ON_LED_Test.jpg |  |  |  |  |

Test Config #3

| Electrode | Polarity LED State / Color (Expected) | Polarity LED State / Color (Actual) | # of Amplitude LEDs Illuminated (Actual) | Polarity LED State / Color (Pass /Fail) |
| --- | --- | --- | --- | --- |
| E0 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E1 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E2 | Illuminated, Red | Illuminated, Red | 5 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E3 | Illuminated, Green | Illuminated, Green | 6 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E4 | Illuminated, Red | Illuminated, Red | 5 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E5 | Illuminated, Green | Illuminated, Green | 5 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E6 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E7 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E8 | Illuminated, Red | Illuminated, Red | 5 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E9 | Illuminated, Green | Illuminated, Green | 5 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E10 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E11 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E12 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E13 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E14 | Illuminated, Red | Illuminated, Red | 5 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E15 | Illuminated, Green | Illuminated, Green | 5 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| Refer RestoreSensorMRI_CP_Stim_ON_LED_Test.jpg |  |  |  |  |

Test Config #4

| Electrode | Polarity LED State / Color (Expected) | Polarity LED State / Color (Actual) | # of Amplitude LEDs Illuminated (Actual) | Polarity LED State / Color (Pass /Fail) |
| --- | --- | --- | --- | --- |
| E0 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E1 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E2 | Illuminated, Green | Illuminated, Green | 8 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E3 | Illuminated, Red | Illuminated, Red | 8 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E4 | Illuminated, Green | Illuminated, Green | 8 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E5 | Illuminated, Red | Illuminated, Red | 8 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E6 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E7 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E8 | Illuminated, Green | Illuminated, Green | 8 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E9 | Illuminated, Red | Illuminated, Red | 8 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E10 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E11 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E12 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E13 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E14 | Illuminated, Green | Illuminated, Green | 8 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E15 | Illuminated, Red | Illuminated, Red | 8 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| Refer PrimeAdvancedMRI_CP_Stim_ON_LED_Test.jpg |  |  |  |  |

Test Config #5

| Electrode | Polarity LED State / Color (Expected) | Polarity LED State / Color (Actual) | # of Amplitude LEDs Illuminated (Actual) | Polarity LED State / Color (Pass /Fail) |
| --- | --- | --- | --- | --- |
| E0 | Illuminated, Red | N/A | N/A | N/A |
| E1 | Not Illuminated, None | N/A | N/A | N/A |
| E2 | Not Illuminated, None | N/A | N/A | N/A |
| E3 | Illuminated, Green | N/A | N/A | N/A |

Test Config #6

| Electrode | Polarity LED State / Color (Expected) | Polarity LED State / Color (Actual) | # of Amplitude LEDs Illuminated (Actual) | Polarity LED State / Color (Pass /Fail) |
| --- | --- | --- | --- | --- |
| E0 | Illuminated, Red | Illuminated, Red | 4 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E1 | Illuminated, Red | Illuminated, Red | 4 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E2 | Illuminated, Red | Illuminated, Red | 4 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E3 | Illuminated, Red | Illuminated, Red | 4 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| Refer Itrel 4B_CP_Stim_ON_LED_Test.jpg |  |  |  |  |

Test Config #7

| Electrode | Polarity LED State / Color (Expected) | Polarity LED State / Color (Actual) | # of Amplitude LEDs Illuminated (Actual) | Polarity LED State / Color (Pass /Fail) |
| --- | --- | --- | --- | --- |
| E0 | Illuminated, Red | Illuminated, Red | 1 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E1 | Illuminated, Green | Illuminated, Green | 1 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E2 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E3 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E4 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E5 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E6 | Illuminated, Red | Illuminated, Red | 1 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E7 | Illuminated, Green | Illuminated, Green | 1 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E8 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E9 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E10 | Illuminated, Red | Illuminated, Red | 1 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E11 | Illuminated, Green | Illuminated, Green | 1 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E12 | Illuminated, Red | Illuminated, Red | 1 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E13 | Illuminated, Green | Illuminated, Green | 1 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E14 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| E15 | Not Illuminated, None | Not Illuminated, None | 0 | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| Refer RestoreAdvancedMRI_CP_Stim_ON_LED_Test.jpg |  |  |  |  |

- Select the Patient Access Screen.  Enable PULSE WIDTH LIMIT and GROUP RATE limit and set to maximum allowable values (without impacting Amplitude limit, which should be 8.70 V based on Test Config 3 stimulation configuration parameters). Press UPDATE and then CLOSE.

- Turn Stimulation OFF.  Record the observed LED Stimulation Box Output below in the correct test config #.  Verify the observed LED Stimulation Box Output matches the expected LED Stimulation Box Output. Indicate Pass or Fail.

| Test Config # | Change stimulation state from ON to OFF (Expected) | Actual | Pass / Fail |
| --- | --- | --- | --- |
| 1 | E0-E15 Polarity LED State / Color is Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated is 0 | E0-E15 Polarity LED State / Color was Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated was 0 Refer RestoreUltraMRI_CP_Stim_OFF_LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 2 | E0-E15 Polarity LED State / Color is Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated is 0 | E0-E15 Polarity LED State / Color was Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated was 0 Refer TestStim_CP_Stim_OFF_LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 3 | E0-E15 Polarity LED State / Color is Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated is 0 | E0-E15 Polarity LED State / Color was Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated was 0 Refer RestoreSensorMRI_CP_Stim_OFF_LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 4 | E0-E15 Polarity LED State / Color is Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated is 0 | E0-E15 Polarity LED State / Color was Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated was 0 Refer PrimeAdvancedMRI_CP_Stim_OFF_LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 5 | E0-E15 Polarity LED State / Color is Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated is 0 | N/A | N/A |
| 6 | E0-E15 Polarity LED State / Color is Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated is 0 | E0-E15 Polarity LED State / Color was Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated was 0 Refer Itrel 4B_CP_Stim_OFF_LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 7 | E0-E15 Polarity LED State / Color is Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated is 0 | E0-E15 Polarity LED State / Color was Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated was 0 Refer RestoreAdvancedMRI_CP_Stim_OFF_LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |

- Select the Electrode Impedance Screen and hit MEASURE ALL IMPEDANCES. Except for Test Config #5, Select 0.70 V amplitude. Select START.

- Record the observed Impedance Classification {Restore Clinician Programmer Application UI} below.  Verify the observed classification matches the expected classification. Indicate Pass or Fail.

| Test Config # | Impedance Classification on Restore Clinician Programmer Application Display (Expected) | Actual | Pass / Fail |
| --- | --- | --- | --- |
| 1 | All Electrodes are classified as “Good” | All Electrodes were classified as “Good” Refer RestoreUltraMRI_ElectordeImpedance_CP.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 2 |  | All Electrodes were classified as “Good” Refer TestStim_ElectordeImpedance_CP.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 3 |  | All Electrodes were classified as “Good” Refer RestoreSensorMRI_ElectordeImpedance_CP.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 4 |  | All Electrodes were classified as “Good” Refer PrimeAdvancedMRI_ElectordeImpedance_CP.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 5 |  | N/A | N/A |
| 6 |  | All Electrodes were classified as “Good” Refer Itrel 4B_ElectordeImpedance_CP.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 7 |  | All Electrodes were classified as “Good” Refer RestoreAdvancedMRI_ElectordeImpedance_CP.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |

- Navigate to the Summary screen and press EXIT WORKFLOW.

- From the splash screen on the Restore Clinician Programmer Application, select REPORTS.

- Ensure “Session Report” is selected within the visible drop-down selection menu.

- Select the most recent session report corresponding to “Regression, #” where # is the current test config, and hit the View Report icon.

- Record the observed report parameters enumerated below in the following two tables.  Verify the observed parameter matches the expected parameter. Indicate Pass or Fail.

  - Patient Name

  - Device Model Name

  - Lead 1 & 2 Model Numbers

  - Program(s) at End Of Session

  - Electrode Impedance

- Close Session Report and the Restore Clinician Programmer Application.  Continue to step 23.

| Test Config # | Name – First Last (Expected) | Device Model (Expected) | Lead Model #’s (Expected) | Name – First Last (Actual) | Device Model (Actual) | Lead Model #’s (Actual) | Name – First, Last (Pass / Fail) | Device Model (Pass / Fail) | Lead Model #’s (Pass / Fail) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 1 Regression | RestoreULTRA / RestoreUltraMRI | 977A1, 977A2 | 1 Regression Refer RestoreUltraMRI_Session_Report.jpg | RestoreUltraMRI Refer RestoreUltraMRI_Session_Report.jpg | 977A1, 977A2 Refer RestoreUltraMRI_Lead_Info.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 2 | 2 Regression | TestStim | 3777, 3877 | 2 Regression Refer TestStim _Session_Report.jpg | TestStim Refer TestStim _Session_Report.jpg | 3777, 3877 Refer TestStim _Lead_Info.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 3 | 3 Regression | RestoreSensorMRI | 3873, 3778 | 3 Regression Refer RestoreSensorMRI _Session_Report.jpg | RestoreSensorMRI Refer RestoreSensorMRI _Session_Report.jpg | 3873, 3778 Refer RestoreSensorMRI _Lead_Info.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 4 | 4 Regression | PrimeADVANCED / PrimeADVANCEDMRI | 39286 | 4 Regression Refer PrimeAdvancedMRI _Session_Report.jpg | PrimeAdvancedMRI Refer PrimeAdvancedMRI _Session_Report.jpg | 39286 Refer PrimeAdvancedMRI _Lead_Info.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 5 | 5 Regression | Restore / RestoreUltraMRI | 3888/3892 / 977A1, 977A2 | N/A | N/A | N/A | N/A | N/A | N/A |
| 6 | 6 Regression | Itrel 4B | 3487A/3890 | 6 Regression Refer Itrel 4B _Session_Report.jpg | Itrel 4B Refer Itrel 4B _Session_Report.jpg | 3487A/3890 Refer Itrel 4B _Lead_Info.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 7 | 7 Regression | RestoreAdvanced MRI | 977A1, 977A2 | 7 Regression Refer RestoreAdvancedMRI _Session_Report.jpg | RestoreAdvancedMRI Refer RestoreAdvancedMRI _Session_Report.jpg | 977A1, 977A2 Refer RestoreAdvancedMRI _Lead_Info.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |

| Test Config # | Program(s) at End of Session (Expected) | Electrode Impedance (Expected) | Program(s) at End of Session (Actual) | Electrode Impedance (Actual) | Program(s) at End of Session (Pass / Fail) | Electrode Impedance (Pass / Fail) |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | A1 is the only Program Cathodes (-): E0, E6, E10, E12 Anodes (+): E1, E7, E11, E13 Amplitude: 2.50 V Pulse Width: 200 μs Rate: 160 Hz | All electrode pair impedance values are >50Ω and <10,000Ω. | A1 was the only Program Cathodes (-): E0, E6, E10, E12 Anodes (+): E1, E7, E11, E13 Amplitude: 2.50 V Pulse Width: 200 μs Rate: 160 Hz Refer RestoreUltraMRI _Programs.jpg | All electrode pair impedance values were >50Ω and <10,000Ω. Refer RestoreUltra MRI_ElectordeImpedance.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 2 | B2 is the only Program Cathodes (-): E1, E7, E11, E13 Anodes (+): E0, E6, E10, E12 Amplitude: 5.00 V Pulse Width: 80 μs Rate: 400 Hz | All electrode pair impedance values are >50Ω and <10,000Ω. | B2 was the only Program Cathodes (-): E1, E7, E11, E13 Anodes (+): E0, E6, E10, E12 Amplitude: 5.00 V Pulse Width: 80 μs Rate: 400 Hz Refer TestStim _Programs.jpg | All electrode pair impedance values were >50Ω and <10,000Ω. Refer TestStim _ ElectordeImpedance.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 3 | C3 is the only Program Present Cathodes (-): E2, E4, E8, E14 Anodes (+): E3, E5, E9, E15 Amplitude: 7.40 V Pulse Width: 90 μs Rate: 800 Hz | All electrode pair impedance values are >50Ω and <10,000Ω. | C3 was the only Program Present Cathodes (-): E2, E4, E8, E14 Anodes (+): E3, E5, E9, E15 Amplitude: 7.40 V Pulse Width: 90 μs Rate: 800 Hz Refer RestoreSensorMRI _Programs.jpg | All electrode pair impedance values were >50Ω and <10,000Ω. Refer RestoreSensorMRI _ ElectordeImpedance.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 4 | C4 is the only Program Present Cathodes (-): E3, E5, E9, E15 Anodes (+): E2, E4, E8, E14 Amplitude: 9.55 V Pulse Width: 120 μs Rate: 80 Hz | All electrode pair impedance values are >50Ω and <10,000Ω. | C4 was the only Program Present Cathodes (-): E3, E5, E9, E15 Anodes (+): E2, E4, E8, E14 Amplitude: 9.55 V Pulse Width: 120 μs Rate: 80 Hz Refer PrimeAdvancedMRI _Programs.jpg | All electrode pair impedance values are >50Ω and <10,000Ω. Refer PrimeAdvancedMRI _ ElectordeImpedance.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 5 | B2 is the only Program Present Cathodes (-): E0 Anodes (+): E3 Amplitude: 10.00 V Pulse Width: 70 μs Rate: 65 Hz | All electrode pair impedance values are >70Ω and <3600Ω. | N/A | N/A | N/A | N/A |
| 6 | A1 is the only Program Present Cathodes (-):E0, E1, E2, E3 Anodes (+):C Amplitude: 6.1 V Pulse Width: 60 μs Rate: 60 Hz | All electrode pair impedance values are >50Ω and <10,000Ω. | A1 was the only Program Present Cathodes (-):E0, E1, E2, E3 Anodes (+): None Amplitude: 6.1 V Pulse Width: 60 μs Rate: 60 Hz Refer Itrel 4B _Programs.jpg | All electrode pair impedance values were >50Ω and <10,000Ω. Refer Itrel 4B _ ElectordeImpedance.jpg | Fail Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 7 | A1 is the only Program Cathodes (-): E0, E6, E10, E12 Anodes (+): E1, E7, E11, E13 Amplitude: 2.50 V Pulse Width: 200 μs Rate: 120 Hz | All electrode pair impedance values are >50Ω and <10,000Ω. | A1 was the only Program Cathodes (-): E0, E6, E10, E12 Anodes (+): E1, E7, E11, E13 Amplitude: 2.50 V Pulse Width: 200 μs Rate: 120 Hz Refer RestoreAdvancedMRI _Programs.jpg | All electrode pair impedance values were >50Ω and <10,000Ω. Refer RestoreAdvancedMRI _ ElectordeImpedance.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |

- Execute this step for Test Config 3, 7 or Test Config 1,4,5  (If MRI devices (Prime Advanced MRI / Restore Ultra MRI) are used), otherwise proceed to next step:

  - Connect to the Neurostimulator with the Restore Clinician Programmer Application. Dismiss any notifications that may appear.

  - Select START under VIEW MRI and go to the MRI screen.

  - Record the observed MRI Eligibility {Restore Clinician Programmer Application UI} in the following table.  Verify the observed MRI Eligibility matches the expected MRI eligibility. Indicate Pass or Fail.

  - Navigate to the Summary screen and press EXIT WORKFLOW.

  - From the splash screen on the Restore Clinician Programmer Application, select REPORTS.

  - Ensure “MRI Report” is selected within the visible drop-down selection menu and select the most recent session report corresponding to “Regression, #” where # is the current test config, and hit the View Report icon.

  - Record the observed MRI Eligibility {MRI Report} in the following table.  Verify the observed MRI Eligibility matches the expected MRI eligibility. Indicate Pass or Fail.

| Test Config # | MRI Eligibility {Restore Clinician Programmer Application UI} (Expected) | MRI Eligibility {MRI Report} (Expected) | MRI Eligibility {Restore Clinician Programmer Application UI} (Actual) | MRI Eligibility {MRI Report} (Actual) | MRI Eligibility Display {UI} (Pass / Fail) | MRI Eligibility {Report} (Pass / Fail) |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Abandoned Leads | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Abandoned Leads | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Abandoned Leads Refer RestoreUltraMRI_MRI_UI.jpg | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Abandoned Leads Refer RestoreUltraMRI_MRI_Report.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 3 | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Lead Models - Lead Tip Location | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Lead Models - Lead Tip Location | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Lead Models - Lead Tip Location Refer RestoreSensorMRI_MRI_UI.jpg | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Lead Models - Lead Tip Location Refer RestoreUltraMRI_MRI_Report.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 4 | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Lead Models - Lead Tip Location - Abandoned Leads | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Lead Models - Lead Tip Location - Abandoned Leads | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Lead Models - Lead Tip Location - Abandoned Leads Refer PrimeAdvancedMRI_MRI_UI.jpg | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Lead Models - Lead Tip Location - Abandoned Leads Refer PrimeAdvancedMRI_MRI_Report.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 5 | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Lead Models - Lead Tip Location - Abandoned Leads | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Lead Models - Lead Tip Location - Abandoned Leads | N/A | N/A | N/A | N/A |
| 7 | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Abandoned Leads | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Abandoned Leads | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Abandoned Leads Refer RestoreAdvancedMRI_MRI_UI.jpg | The neurostimulator system MRI scan-type eligibility cannot be determined. Not eligible for full body scan due to: - Abandoned Leads Refer RestoreAdvancedMRI_MRI_Report.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |

- Place the MyStim Patient Programmer over the Neurostimulator and press the Sync Button .

- Turn Stimulation ON by hitting the ON button  .  Record the observed LED Stimulation Box Output in the following table.  Verify the observed LED Stimulation Box Output matches the expected LED Stimulation Box Output. Indicate Pass or Fail.

| Test Config # | After changing stimulation state from OFF to ON utilizing MyStim (Expected) | After changing stimulation state from OFF to ON utilizing MyStim (Actual) | Pass / Fail |
| --- | --- | --- | --- |
| 1 | Polarity LED State / Color and # of Amplitude LEDs Illuminated matches what was recorded in Step 12. | Polarity LED State / Color and # of Amplitude LEDs Illuminated matches what was recorded in Step 12. Refer RestoreUltraMRI_MyStim_Stim_ON_LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 2 |  | Polarity LED State / Color and # of Amplitude LEDs Illuminated matches what was recorded in Step 12. Refer TestStim_MyStim_Stim_ON_LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 3 |  | Polarity LED State / Color and # of Amplitude LEDs Illuminated matches what was recorded in Step 12. Refer RestoreSensorMRI_MyStim_Stim_ON_LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 4 |  | Polarity LED State / Color and # of Amplitude LEDs Illuminated matches what was recorded in Step 12. Refer PrimeAdvancedMRI_MyStim_Stim_ON_LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 5 |  | N/A | N/A |
| 6 |  | Polarity LED State / Color and # of Amplitude LEDs Illuminated matches what was recorded in Step 12. Refer Itrel 4B_MyStim_Stim_ON_LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 7 |  | Polarity LED State / Color and # of Amplitude LEDs Illuminated matches what was recorded in Step 12. Refer RestoreAdvancedMRI_MyStim_Stim_ON_LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |

- Attempt to increment Amplitude by pressing the INCREASE button .  Attempt to decrement Amplitude by pressing the DECREASE button . Record the observed behavior in the following table.  Verify the observed Amplitude Adjustment behavior matches the expected Amplitude Adjustment behavior. Indicate Pass or Fail.

- Use the arrow right to highlight Pulse Width. Attempt to increment Pulse Width by pressing the INCREASE button .  Attempt to decrement Pulse Width by pressing the DECREASE button . Record the observed behavior in the following table.  Verify the observed Pulse Width Adjustment behavior matches the expected Pulse Width Adjustment behavior. Indicate Pass or Fail.

- Press right on the directional pad to highlight Rate. Attempt to increment Rate by pressing the INCREASE button .  Attempt to decrement Rate by pressing the DECREASE button . Record the observed behavior in the following table.  Verify the observed Rate Adjustment behavior matches the expected Rate Adjustment behavior. Indicate Pass or Fail.

| Test Config # | Adjust Amplitude (Expected) | Adjust Pulse Width (Expected) | Adjust Rate (Expected) | Adjust Amplitude (Actual) | Adjust Pulse Width (Actual) | Adjust Rate (Expected) | Adjust Amplitude (Pass / Fail) | Adjust Pulse Width (Pass / Fail) | Adjust Rate (Pass / Fail) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | The MyStim Patient Programmer allows amplitude to be incremented and decremented. | The MyStim Patient Programmer allows pulse width to be incremented and decremented. | The MyStim Patient Programmer allows rate to be incremented and decremented. | The MyStim Patient Programmer allowed amplitude to be incremented and decremented. Refer RestoreUltraMRI_Amplitue_Increment.jpg Refer RestoreUltraMRI_Amplitue_Decrement.jpg | The MyStim Patient Programmer allowed pulse width to be incremented and decremented. Refer RestoreUltraMRI_PulseWidth_Increment.jpg Refer RestoreUltraMRI_PulseWidth_Decrement.jpg | The MyStim Patient Programmer allowed rate to be incremented and decremented. Refer RestoreUltraMRI_Rate_Increment.jpg Refer RestoreUltraMRI_Rate_Decrement.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 2 |  |  |  | The MyStim Patient Programmer allowed amplitude to be incremented and decremented. Refer TestStim_Amplitue_Increment.jpg Refer TestStim_Amplitue_Decrement.jpg | The MyStim Patient Programmer allowed pulse width to be incremented and decremented. Refer TestStim_PulseWidth_Increment.jpg Refer TestStim_PulseWidth_Decrement.jpg | The MyStim Patient Programmer allowed rate to be incremented and decremented. Refer TestStim_Rate_Increment.jpg Refer TestStim_Rate_Decrement.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 3 |  |  |  | The MyStim Patient Programmer allowed amplitude to be incremented and decremented. Refer RestoreSensorMRI_Amplitue_Increment.jpg Refer RestoreSensorMRI _Amplitue_Decrement.jpg | The MyStim Patient Programmer allowed pulse width to be incremented and decremented. Refer RestoreSensorMRI _PulseWidth_Increment.jpg Refer RestoreSensorMRI _PulseWidth_Decrement.jpg | The MyStim Patient Programmer allowed rate to be incremented and decremented. Refer RestoreSensorMRI _Rate_Increment.jpg Refer RestoreSensorMRI _Rate_Decrement.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 4 |  |  |  | The MyStim Patient Programmer allowed amplitude to be incremented and decremented. Refer PrimeAdvancedMRI_Amplitue_Increment.jpg Refer PrimeAdvancedMRI _Amplitue_Decrement.jpg | The MyStim Patient Programmer allowed pulse width to be incremented and decremented. Refer PrimeAdvancedMRI _PulseWidth_Increment.jpg Refer PrimeAdvancedMRI _PulseWidth_Decrement.jpg | The MyStim Patient Programmer allowed rate to be incremented and decremented. Refer PrimeAdvancedMRI _Rate_Increment.jpg Refer PrimeAdvancedMRI _Rate_Decrement.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 5 |  |  |  | N/A | N/A | N/A | N/A | N/A | N/A |
| 6 |  |  |  | The MyStim Patient Programmer allowed amplitude to be incremented and decremented. Refer Itrel 4B_Amplitue_Increment.jpg Refer Itrel 4B_Amplitue_Decrement.jpg | The MyStim Patient Programmer allowed pulse width to be incremented and decremented. Refer Itrel 4B_PulseWidth_Increment.jpg Refer Itrel 4Btim_PulseWidth_Decrement.jpg | The MyStim Patient Programmer allowed rate to be incremented and decremented. Refer Itrel 4B_Rate_Increment.jpg Refer Itrel 4B_Rate_Decrement.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 7 |  |  |  | The MyStim Patient Programmer allowed amplitude to be incremented and decremented. Refer RestoreAdvancedMRI_Amplitue_Increment.jpg Refer RestoreAdvancedMRI _Amplitue_Decrement.jpg | The MyStim Patient Programmer allowed pulse width to be incremented and decremented. Refer RestoreAdvancedMRI _PulseWidth_Increment.jpg Refer RestoreAdvancedMRI _PulseWidth_Decrement.jpg | The MyStim Patient Programmer allowed rate to be incremented and decremented. Refer RestoreAdvancedMRI _Rate_Increment.jpg Refer RestoreAdvancedMRI _Rate_Decrement.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. | Pass Electronically Signed in MRCS 614848, 11 December 2024. |

- Turn Stimulation OFF by hitting the OFF button .  Record the observed LED Stimulation Box Output in the following table.  Verify the observed LED Stimulation Box Output matches the expected LED Stimulation Box Output. Indicate Pass or Fail.

| Test Config # | After changing stimulation state from ON to OFF utilizing MyStim (Expected) | After changing stimulation state from ON to OFF utilizing MyStim (Expected) | Pass / Fail |
| --- | --- | --- | --- |
| 1 | E0-E15 Polarity LED State / Color is Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated is 0 | E0-E15 Polarity LED State / Color was Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated is 0 Refer RestoreUltraMRI_MyStim_Stim_OFF_LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 2 |  | E0-E15 Polarity LED State / Color was Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated is 0 Refer TestStim_MyStim_Stim_OFF_LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 3 |  | E0-E15 Polarity LED State / Color was Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated is 0 Refer RestoreSensorMRI_MyStim_Stim_OFF_LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 4 |  | E0-E15 Polarity LED State / Color was Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated is 0. Refer PrimeAdvancedMRI_MyStim_Stim_OFF_LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 5 |  | N/A | N/A |
| 6 |  | E0-E15 Polarity LED State / Color was Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated is 0 Refer Itrel 4B_MyStim_Stim_ OFF _LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |
| 7 |  | E0-E15 Polarity LED State / Color was Not Illuminated / None E0-E15 # of Amplitude LEDs Illuminated is 0. Refer RestoreAdvancedMRI_MyStim_Stim_OFF_LED_Test.jpg | Pass Electronically Signed in MRCS 614848, 11 December 2024. |

### Restore Testing – Set Neurostimulators to Shelf State

The device may be set to Shelf State utilizing the Lab Programmer.

- Connect the Lab Programmer Telemetry Head to the PC with Lab Programmer installed.

- Open “Rx1 ProgApp Validated” application.

- Once the application opens, select “Launch” button from the lower right-hand corner of the “Rx1 ProgApp Validated” window to connect to the “NS1 Test Platform” executable (this will launch in the background, no interaction should be needed given Lab Programmer is already installed on the machine).

- After observing the “NS1 Test Platform” executable open in the background, select the “Connect” button from the lower right-hand corner of the “Rx1 ProgApp Validated” window and wait approximately 5 seconds for setup to complete (the last transaction should be displayed in the “Raw Received Data” box as “SIO:1003:100201:F956”.

- Select appropriate “Device Family” from drop-down menu.

| Device | Device Family Selection |
| --- | --- |
| 37711, Restore / 37713, RestoreAdvanced / 97713, RestoreAdvancedMRI | 0101 Restore/RestoreAdvanced |
| 37701, RestorePrime / 37702, PrimeAdvanced / 97702, PrimeAdvancedMRI | 0102 RestorePrime/ PrimeAdvanced |
| 37712, RestoreUltra / 97712, RestoreUltraMRI / 37022, Test Stim ENS | 0103 Ultra/Activa RC |
| 37703, Itrel4A / 37704, Itrel4B | 0104 Activa PC / Activa SC |
| 37714, RestoreSensor / 97714 RestoreSensorMRI | 0105 Restore Sensor |

- Position the telemetry head directly over the device (Lab Programmer Telemetry head LED should face ceiling; The etched Medtronic logo should also face the ceiling).

- Select “ID Request” from the upper left-hand corner of the ”Rx1 ProgApp Validated” window and wait for “Device ID” fields to populate (approximately 2-5 seconds).

If successful, the “Device ID” fields should look similar to the following:

If unsuccessful, the “Device ID” fields may look like the following:

In the event it is unsuccessful, simply repeat this step again OR close both application windows and return to the first step under “Setting the INS to Shelf State”.

- Navigate to the “Other” tab within the “Rx1 Lab Programmer Validated” window.

- Select the “Reset ENS or Rechargeable INS” button located towards the center of the “Rx1 Lab Programmer Validated” window for rechargeable devices or ENS. Select Reset Primary Cell INS for primary cell devices.

- Wait for Reset to complete (It may take a few minutes).

- Device Reset is complete once you see a message similar to the one shown below:

Close “Rx1 Lab Programmer Validated” window. Close “NS1 Test Platform” window